SA6 mini频谱分析仪&信号发生器

# 4.4G Spectrum Analyzer & Signal Genertor

产品说明书

Product Manual

Product Manual

Packing

1.  4.4G Spectrum Analyzer & Signal Genertor       x1

2.  Connecting cable USB Type-C                x1

3.  SMA-kk  female connector                  x1

4.  Touch pen                                 x1

Panel infomation

- The output of the signal generator TG OUT;

- Antenna input RF IN;

- Housing;

- Colour resistive screen of 3.2” ;

- OFF/ON;

- Battery charge indicator charge;

- Connector type C USB;

- Standby button on/off .

9.«MKR» setting markers. Long (more than 2 seconds) button holding opens the menu of markers and peak values of the analyzed signal.

10.«FREQ»-displays the frequency parameters on the screen. Long press (more than 2 seconds) to open the menu for setting the viewing frequency range.

11. Buttons to move the screen scale grid and change the frequency.

12.«SPAN»(overview frequency range) -buttons " -" and " + " change the frequency range of the analyzed signal.

13.«AMPL»-output amplitude parameters on the screen. Long (more than 2 seconds) holding the button opens the menu for setting the amplitude parameters.

14.«MENU»-the button opens the main menu of the device.

Specifications

| Parameter name |  | Value |
| --- | --- | --- |
| Displayed frequency range |  | 35-4400 MHz |
| Measured frequency range |  | 45-4400 MHz |
| Dynamic range | for 45-3000 MHz band | 75 dB |
|  | for 3000-4400 MHz band | 70 dB |
| Maximum bandwidth |  | 4355 MHz |
| IF bandwidth (fixed) |  | 200 kHz |
| Sweep speed with a span of more than 100 MHz |  | 600 MHz/s |
| Sweep time of 100 MHz span |  | < 0.25 s |
| Detectable signal duration at 10 MHz scan span |  | 20 ms |
| Noise floor | in the band up to 3000 MHz | < -100 dBm |
|  | in the band 3000-4500 MHz | < -95 dBm |
|  | in the band 4500-6200 MHz | < -90 dBm |
| Input impedance |  | 50 Ohm |
| SWR in the operating frequency range |  | < 1,5 |
| Internal attenuator |  | 0-30 dB |
| Error of attenuation | up to 10 dB | 2 dB |
|  | from 10 to 20 dB | 5 dB |
|  | from 20 to 30 dB | 8 dB |
| Frequency range of built-in tracking generator |  | 35-6200 MHz |
| Tracking generator output power |  | -15 .. -25 dBm |
| Accuracy of display of a signal within a dynamic range |  | 2 dB |
| Maximum input power | with 0 dB attenuator | + 10 dBm |
|  | with attenuator > 20 dB | + 20 dBm |
| The measured maximum input signal |  | + 10 dBm |
| Maximum RF input DC voltage |  | 5V |
| Maximum supply current | when running on battery | 450 mA |
|  | when running on USB (in charging mode) | 500 mA |
| Battery capacity |  | 3000 mAh |
| Continuous operation time from accumulator |  | 3 h |
| Battery charging time |  | ~4 h |
| Dimensions |  | 155×81×27 mm |
| Weight |  | 0.49kg |

- Notice：We are not responsible for the damage of the machine caused by privately upgrading the software version.